# EIS analysis report

Model: Randles (Rs + Rct‖Cdl)

Data file: randles_Rs10_Rct30_C100mF.csv

Generated on 2026-10-12 13:23

# 1. Equivalent-circuit fit

Electrical parameters of the chosen model, found by least-squares fitting. They describe the circuit, not yet the chemistry.

Rs:  10 Ω

Rct:  29.98 Ω

Cdl:  0.0001002 F

RMSE (complex):  0.06258 Ω

χ² (σ = 1):  3.9

# 2. Measured EIS features (no model)

Read directly from the data. If the system behaves like a Randles circuit: Rhigh ≈ Rs, Rlow - Rhigh ≈ Rct, and fapex is related to Rct and Cdl as below.

Rhigh:  10.11 Ω

Rlow:  39.89 Ω

Rlow - Rhigh:  29.78 Ω

fapex:  52.11 Hz

LaTeX:  f_{apex} = \frac{1}{2 \pi \, R_{ct} \, C_{dl}}

# 3. Electrochemical interpretation

Derived quantities. Each one is valid only under the assumptions listed with it.

## Inputs provided

Temperature: 25 °C

Electrode area: 1 cm²

Number of electrons n: 1

Electrode spacing L: 1 cm

Electrolyte cross-section: 1 cm²

## Shared assumptions

A1: The fitted circuit is Rs + (Rct ‖ Cdl): one semicircle, and Cdl is an ideal capacitor (no CPE / non-ideal behaviour).

## Time constant τ

Value:  0.003004 s

LaTeX:  \tau = R_{ct} \, C_{dl}

Relies on shared assumption(s): A1

Notes:

Describes how fast the interface charges/discharges. It is a property of the fit, not yet a chemical statement.

## Characteristic frequency fc

Value:  52.98 Hz

LaTeX:  f_c = \frac{1}{2 \pi \tau}

Relies on shared assumption(s): A1

Notes:

Should match the apex frequency measured directly from the Nyquist plot (fapex); a large mismatch means the single-semicircle model does not describe the data well.

## Specific double-layer capacitance

Value:  0.0001002 F/cm2

LaTeX:  C_{spec} = \frac{C_{dl}}{A}

Relies on shared assumption(s): A1

Additional assumptions:

A is the true electroactive area, not just the geometric area.

The fitted capacitance is really the double layer (not adsorption pseudocapacitance or a coating).

Notes:

100 µ F/cm2 is outside the usual ~5-100 µ F/cm2 range for flat electrodes in aqueous electrolyte. Possible reasons: real area differs from the area you entered (roughness, porosity), a surface film or functionalised layer, a non-aqueous electrolyte, or a non-ideal (CPE) interface. EIS alone cannot tell which.

## Exchange current density j0

Value:  0.0008569 A/cm2

LaTeX:  j_0 = \frac{R\,T}{n\,F\,R_{ct}\,A}

Relies on shared assumption(s): A1

Additional assumptions:

Measured at equilibrium (open-circuit) with both the oxidised and reduced species present, small AC amplitude (linearised Butler-Volmer).

Rct is purely the charge-transfer resistance: no diffusion (Warburg) or film resistance mixed into it.

Single, simple electron-transfer reaction; n is the number of electrons of that reaction (ambiguous for multi-step mechanisms).

A is the true electroactive area.

T = 298.15 K.

Notes:

If Rct also contains other resistive contributions, j0 is underestimated.

## Electrolyte conductivity

Value:  0.09999 S/cm

LaTeX:  \kappa = \frac{K}{R_s}, \quad K = \frac{L}{A}

Additional assumptions:

Rs is entirely electrolyte resistance (wires, contacts and reference-electrode junction add to it).

Uniform current distribution between electrodes, so that the cell constant K = L/A is valid.

Rs is read at high frequency, where the cell is purely resistive.

Notes:

1 S/cm = 1000 mS/cm. Compare to a calibrated standard solution to validate the cell constant.

## Consistency check: fit vs data

fc (fit):  52.98 Hz

fapex (data):  52.11 Hz

Difference:  1.7 %

The fit and the data agree on the time constant. This supports the single-semicircle model but does not prove it.

# 4. Figures

| Nyquist plot (data) | Bode plot (data) |
| --- | --- |

| Nyquist: data vs fitted model | Bode: data vs fitted model |
| --- | --- |